Report

Stack Testing Measurements

in flue gas of the

Fábrica de Cimento, Secil Outão, Portugal

Kiln 8 and 9

September 2001

Accreditation by the German Authorities (Notification) related to §§ 26, 28 BImSchG, 2. BImSchV, §§ 26,28 of 13.BImSchV, § 10 of 17. BImSchV, Nr. 3.2 TA Luft

Emission- and Ambient Air Measurement, Olfactometry, Calibration and Function Test

Geierstraße 1			Bankinformation:				Board members:

22305 Hamburg, Germany		Commerzbank Hamburg				Dr. Michael Ball, Olaf Päpke

Phone (+4940) 69 70 960		BLZ 200 400 00					HRB 22799

Fax (+4940) 69 70 9699			Account No. 270 78 26				Amtsgericht Hamburg

e-mail: ergo.umwelt@t-online.de

Report

Stack Testing Measurements

in flue gas of the

Fábrica de Cimento, Secil Outão, Portugal

Kiln 8 and 9

September 2001

| Ordered by: | Companhia Geral de Cal e Cimento, S.A., SECIL, Outão, Portugal |
| --- | --- |
| Order n° | 4500035986 |
| Date of order: | 01-04-02 |
| Location of measurement | Fábrica SECIL Outão Apartado 71 2901-901 Setúbal |
| Part of plant | Kiln 8 – 9 |
| Period of measurement: | 01-04-19 – 01-04-25 |
| ERGO project number: | A-179-01-200-SECIL4 Outao-b.doc |
| Project manager: | Dr. U. Düwel, Dipl.-Ing. Schröder |
| Phone: | (+4940) 69 70 96 – 0 |
| Number of pages: | 7 |

Contents										       page

1	Introduction	4

2	Plant characterization	4

3	Summarized test results	5

4	Characterization of measurements	7

Annex A: Technical report

Annex B: Operating conditions

Introduction

The Companhia Geral de Cal e Cimento, S.A., Portugal, SECIL, instructed ERGO Forschungsgesellschaft mbH, Hamburg, Germany, to perform measurements with regard to concentrations of various components in the flue gas of two cement kilns, located at the Outão plant in Portugal. The tests were carried out at kiln no. 8 (stack no. J8P11) and kiln no. 9 (stack no. J9P11A8 and J9P11A9) to determine the concentrations of dioxins/ furans (polychlorinated-p-dibenzodioxins and polychlorinated-p-dibenzofurans), heavy metals, hydrogen chloride, hydrogen fluoride, sulphur dioxide, nitrogen oxides expressed as nitrogen dioxide, carbon monoxide, dust, and total organic carbon (TOC) in accordance with the EC – Council Directive 2000/76/EC of December 4th, 2000 (Incineration of Waste).

The tests aim to show the emission characteristics of the investigated kilns which are equipped with an electrostatic precipitator as a flue gas cleaning facility, in order to be able to compare the results with future installations.

Plant characterization

The cement plant can be described as a four-stage cyclone preheater kiln equipped with planetary coolers.

The location and sampling period of the investigations are listed in the following table:

Table 2.1: Location and date of measurements

| Company | Plant location | Kiln #/ stack # | Measuring period |
| --- | --- | --- | --- |
| SECIL | Outão | 8 (J8p11) | April 19th – April 20th |
| SECIL | Outão | 9 (J9P11A8 and J9P11A9) | April 24th – April 25th |

During the stack testing measurements, the cleaning of the flue gas was carried out using the electrostatic precipitator (conventional system). Fabric filters were still under construction at this time and thus not applied. The raw mill was in operation during all tests.

A short documentation of the key data concerning operating conditions was compiled from recordings from the plant process data system and placed at our disposal. The data listed in the following tables show normal operating conditions.

Table 2.2: Operating conditions of kiln #8

| Kiln 8 |
| --- |
| Date | Period | Raw mill feed [t/h] | Kiln feed [t/h] | Clinker production [t/h] |
| 19/04/2001 | 08:00-13:00 | 161,0 | 171,8 | 95,4 |
| 19/04/2001 | 13:00-20:00 | 207,0 | 174,6 | 97,0 |
| 20/04/2001 | 08:00-13:00 | 200,4 | 168,6 | 93,8 |
| 20/04/2001 | 13:00-20:00 | 204,9 | 173,6 | 96,3 |

Table 2.3: Operating conditions of kiln #9

| Kiln 9 |
| --- |
| Date | Period | Raw mill feed [t/h] | Kiln feed [t/h] | Clinker production [t/h] |
| 24/04/2001 | 08:00-13:00 | 298,4 | 245,2 | 143,6 |
| 24/04/2001 | 13:00-20:00 | 297,3 | 255,3 | 149,5 |
| 25/04/2001 | 08:00-13:00 | 316,0 | 260,7 | 152,8 |
| 25/04/2001 | 13:00-20:00 | 299,4 | 260,0 | 152,3 |

- Summarized test results

The measurement results are summarized in the following tables.

In accordance with the EC – Council Directive 2000/76/EC of December, 4th, 2000 the concentrations are standardized to a  temperature of  273 K , a  pressure of  1013 hPa of the dry flue gas and related to an oxygen content (dry) of 10 Vol%.

Table 3.:	SECIL, Outão, kiln #8

| Components | Unit | 19.04.01 | 20.04.01 |
| --- | --- | --- | --- |
| Carbon monoxide (CO) (1) | [mg/m³] | 635 | 458 |
| Carbon monoxide (CO) (2) | [mg/m³] | 632 | 438 |
| Nitrogen oxide expr. as NO2 (1) | [mg/m³] | 1020 | 631 |
| Nitrogen oxide expr. as NO2 (2) | [mg/m³] | 1080 | 1000 |
| Total organic carbon (1) | [mg/m³] | 30 | - (2) |
| Total organic carbon (2) | [mg/m³] | 26 | - (2) |
| Hydrogen chloride (HCl) | [mg/m³] | 1,7 | 1,7 |
| Hydrogen fluoride (HF) | [mg/m³] | 0,4 | 0,3 |
| Sulphur dioxide (SO2) | [mg/m³] | 39 | 6,3 |
| Total dust | [mg/m³] | 5,7 | -(2) |
| Sum of cadmium (Cd) and thallium (Tl) | [mg/m³] | 0,0061 | < 0,005 |
| Mercury (Hg) | [mg/m³] | < 0,005 | < 0,005 |
| Sum of Sb to V | [mg/m³] | 0,061 | < 0,05 |
| Dioxins / furans, I-TEQ-value EN 1948 | [ng/m³] | < 0,01 | < 0,01 |

1 antimony (Sb), arsenic (As), lead (Pb), chromium (Cr), cobalt (Co), copper (Cu), manganese (Mn), nickel (Ni), vanadium (V)

2 evaluation not possible

Table 3.:	SECIL, Outão, kiln #9, stack no. J9P11A8

| Components | Unit | 24.04.01 | 25.04.01 |
| --- | --- | --- | --- |
| Carbon monoxide (CO) (1) | [mg/m³] | 613 | 520 |
| Carbon monoxide (CO) (2) | [mg/m³] | 581 | 556 |
| Nitrogen oxide expr. as NO2 (1) | [mg/m³] | 785 | 825 |
| Nitrogen oxide expr. as NO2 (2) | [mg/m³] | 768 | 810 |
| Total organic carbon (1) | [mg/m³] | -(2) | 23 |
| Total organic carbon (2) | [mg/m³] | -(2) | 24 |
| Hydrogen chloride (HCl) | [mg/m³] | 1,0 | 0,8 |
| Hydrogen fluoride (HF) | [mg/m³] | 0,2 | 0,3 |
| Sulphur dioxide (SO2) | [mg/m³] | 83 | 88 |
| Total dust | [mg/m³] | 16 | 13 |
| Sum of cadmium (Cd) and thallium (Tl) | [mg/m³] | < 0,005 | < 0,005 |
| Mercury (Hg) | [mg/m³] | < 0,005 | < 0,005 |
| Sum of Sb to V 1 | [mg/m³] | < 0,05 | < 0,05 |
| Dioxins / furans, I-TEQ-value EN 1948 | [ng/m³] | < 0,01 | < 0,01 |

Table 3.:	SECIL, Outão, kiln #9, stack no. J9P11A9

| Components | Unit | 24.04.01 | 25.04.01 |
| --- | --- | --- | --- |
| Carbon monoxide (CO) (1) | [mg/m³] | 682 | 587 |
| Carbon monoxide (CO) (2) | [mg/m³] | 744 | 600 |
| Nitrogen oxide expr. as NO2 (1) | [mg/m³] | 1150 | 721 |
| Nitrogen oxide expr. as NO2 (2) | [mg/m³] | 1120 | 727 |
| Total organic carbon (1) | [mg/m³] | -2 | 22 |
| Total organic carbon (2) | [mg/m³] | -2 | 22 |
| Hydrogen chloride (HCl) | [mg/m³] | 0,9 | 0,5 |
| Hydrogen fluoride (HF) | [mg/m³] | 0,2 | 0,2 |
| Sulphur dioxide (SO2) | [mg/m³] | 105 | 68 |
| Total dust | [mg/m³] | 20 | 15 |
| Sum of cadmium (Cd) and thallium (Tl) | [mg/m³] | < 0,005 | < 0,005 |
| Mercury (Hg) | [mg/m³] | < 0,005 | < 0,005 |
| Sum of Sb to V 1 | [mg/m³] | < 0,05 | < 0,05 |
| Dioxins / furans, I-TEQ-value EN 1948 | [ng/m³] | < 0,01 | < 0,01 |

1 antimony (Sb), arsenic (As), lead (Pb), chromium (Cr), cobalt (Co), copper (Cu), manganese (Mn), nickel (Ni), vanadium (V)

2 evaluation not possible

The concentration levels of dioxins/ furans and total organic carbon are in the same order of magnitude as those ascertained at German cement plants. The concentration of heavy metals, hydrogen chloride, hydrogen fluoride, sulfur dioxide, carbon monoxide and nitrogen oxide is determined by the content of these compounds in the raw material, respectively by operating conditions.

As far as the consecutively measured components are concerned, fluctuations in measurement values due to varying plant conditions are conceivable.

Characterization of measurements

Each of the congeners listed in Annex I of the Council Directive 2000/76/EC of December 4th, 2000, is determined for the measurement of dioxins/ furans. The I-TEQ value is evaluated following the procedure described in Annex 1. This procedure also complies with the European Standard EN 1948.

The heavy metal values include particle bounded forms, the gaseous (vapor) forms and their compounds in accordance with the Council Directive 2000/76/EC. Tables 3.1 to 3.3 show the sum of groups of elements related to Annex II.1.1 of the Council Directive. In case that the result for the single element was below the detection limit half of the value of the detection limit was set into the sum up calculation.

Dioxins/ furans and heavy metals are detectable even at very low concentrations through use of laboratory procedures based on instrumental analysis. The detection limit, the uncertainty and the blank levels of the measuring procedures - respectively small inevitable process variations - cause a specific range which cannot be used for comparative tests. Therefore, levels within this range are not reported but indicated by “lower than”.

For the components hydrogen chloride, hydrogen fluoride, sulfur dioxide, carbon monoxide, nitrogen oxides and total organic carbon (TOC), the resultant values show averages over ½ hour measurement period. The values for dioxins/ furans are based on an average sampling time of six hours, the values for total dust and the heavy metals on an average time of one hour.

The flow rate of the sampled gas stream is adjusted to isokinetic conditions on the sampling nozzle during sampling of dioxins, dust and heavy metals.

The sampling of dioxins / furans, heavy metals and dust was carried out over the cross-section of the stacks following ISO 9096. For kiln #9, simultaneous measurements for these components were carried out at both stacks.

Technical details of the measurements are reported in Annex A, those of the operating conditions in Annex B.

ERGO Forschungsgesellschaft mbH

Hamburg, September 10th, 2001

Dr. Uwe Düwel					Dipl. Ing. Hauke Schröder